## S-Series Performance Report
Transmitter Performance Report
Calculations based on published performance specifications by Schneider-Electric, updated October 28, 2020.
Application Data |  | Pressure Transmitter Accuracy Results
Pressure Format | Differential | Upper Range Limit (URL) | 200.0001 | inH2O
Units | inH2O | Min Span (URV - LRV) | 0.5 | inH2O
Upper Range Value (URV) | 200 | Transmitter Turndown | 1
Lower Range Value (LRV) | 0 | Reading Accuracy Down To | 66.6667 | inH2O
Point of Interest (POI) | 2 | FoxCal | FoxCal On
Maximum Process Pressure | 250 | Ref. Error at URV | 0.1 | inH2O
Transmitter Series | IDP10S | Ref. Error at LRV | 0.033333 | inH2O
Body Style | Biplanar or Low-Profile | Ref. Error at POI | 0.033333 | inH2O
Transmitter Range | B | Installed +/- Error at URV | 0.12 | inH2O
Selection Compatibility | Acceptable | Installed +/- Error at LRV | 0.053333 | inH2O
Application | dP Flow | Installed +/- Error at POI | 0.053333 | inH2O
Output Type | Analog (4-20mA) | Max Working Pressure | 100467.7576 | inH2O
|  | Proof Pressure | 401760.2 | inH2O
Installation Factors |  | dP Flow Application
Ambient Temperature Effect | No Change | Max Flow Rate (at URV) | 500 | GPM
Time in Service (years) | 0 | dP Element Error (% max flow) | 0.6 | %
Static Pressure Effect | None | Flow Rate at POI | 50 | GPM
Supply Voltage Effect (VDC) | 0 | Flow +/- Error at URV | 3.0037 | GPM
Vibration Effect | 0 | Flow +/- Error at LRV | 8.6987 | GPM
RFI Effect | None | Flow +/- Error at POI | 3.0722 | GPM
Lightening or Surge Events | 0 | Totalized Daily +/- Error at URV | 4,325 | Gal
Shunt Resistor Tolerance (%) | 0 | Totalized Daily +/- Error at LRV | 12,530 | Gal
Input Module Error (%) | 0 | Totalized Daily +/- Error at POI | 4,424 | Gal
www.nealsystems.com ¦ 122 Terry Drive, Newtown, PA 18940 ¦ sales@nealsystems.com ¦ 215-968-7577
Note on C29: Select +/-50 deg F for applications subject to ambient temperature fluctuation, such as outdoor installations.
Note on C30: Enter the number of years between the factory calibration and the measurment accuracy calculation. Maximum of 10 years.
Note on C31: For differential pressure applications, select Up to 1000 PSI if the static process pressure is high.
Note on C32: Enter the expected supply voltage variation in Volts DC.
Note on C33: Enter the expected installation vibration level in g's peak to peak. The aluminum housing is rated to 2g, The SS housing is rated to 3g.
Note on C34: Select High for applications subject to radio frequency interference in the range of 30 to 1000 MHz and field intensity up 30 V/m.
Note on C35: Enter a whole number between 0 and 10, representing the number of lightning or surge events up 2000V.
Note on C36: For analog output applications, enter the tolerance of the external shunt/resistor installed in the circuit as a percent. This is usually specified by the appropriate color band on the resistor itself. If using an analog input module with an internal resistor, enter 0 in this field and use the next field to account for the total error of the input measurement. This field is ignored for digital applications.
Note on C37: For analog applications, enter the percent error of the PLC, data recorder, or I/O module used to read the 4-20mA output generated by the transmitter. This field is ignored for digital applications.

## B1 (hidden)
Characteristic | Status | Reference | Calc | Status | Lists | [6 empty] | Status | Model Var | Model | URL or Span Limit | Min Span | Lower Range Min | Reading Error Turndown | Reading Error % | Span Reference Error % | Reference Turndown | URL/Span Constant | Max Overrange | Max Proof | Temp Effect URL | Temp Effect Span | Temp Error Units | Stability/YR URL | Stability Error Units | Static Pressure Span | Static P URV Error Units | Static P LRV Error Units | Static P POI Error Units | Supply Power/V Span | Supply Voltage Error Units | Vibration URL | Vibration Error Units | RFI Span | RFI Error Units | Surge Span | Surge Error Units | Analog Output Error %Span | Analog Output Error Units | Resistor Error Units | Analog Input Error Units
Absolute | 0 | Units Conv. | 27.7076 |  | Format | Absolute | Gauge | Differential |  |  |  | 0 | IAP05S, Dir, D | IAP05S | 200 | 0.5 | 0 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 300 | 800 | 0.03 | 0.06 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
Gauge | 0 | URV in PSI | 7.2182 |  | Units | PSI | BAR | inH2O |  |  |  | 0 | IAP05S, Dir, E | IAP05S | 2000 | 5 | 0 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 3000 | 8000 | 0.03 | 0.06 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
Differential | 1 | LRV in PSI | 0 |  | Absolute Series | IAP05S | IAP10S | IAP50S |  |  |  | 0 | IAP05S, Bi, B | IAP05S | 7.2182 | 0.018046 | 0 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
PSI | 0 | POI in PSI | 0.072182 |  | Gauge Series | IGP05S | IGP10S | IGP50S |  |  |  | 0 | IAP05S, Bi, C | IAP05S | 36.0912 | 0.089867 | 0 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
BAR | 0 | Max Pressure in PSI | 9.0228 |  | Differential Series | IDP05S | IDP10S | IDP50S |  |  |  | 0 | IAP05S, Bi, D | IAP05S | 300 | 3.75 | 0 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
inH2O | 1 | Max Range | 7.2182 |  | AP/GP Body Style | Direct Connect | Biplanar or Low-Profile |  |  |  |  | 0 | IAP05S, Bi, E | IAP05S | 3000 | 37.5 | 0 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.08 | 0.25 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
05S | 0 | Min Range | 0 |  | dP Body Style | Biplanar or Low-Profile | [5 empty] | 0 | IGP05S, Dir, D | IGP05S | 200 | 0.5 | 0 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 300 | 800 | 0.03 | 0.06 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
10S | 1 | Span in PSI | 7.2182 | 0 | 05S AP Direct Spans | D | E |  |  |  |  | 0 | IGP05S, Dir, E | IGP05S | 2000 | 0.5 | 0 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 3000 | 8000 | 0.03 | 0.06 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
50S | 0 | Span Range Index | 10 | 0 | 05S AP Biplanar Spans | B | C | D | E |  |  | 0 | IGP05S, Dir, F | IGP05S | 6000 | 75 | 0 | 1 | 0.075 | 0.075 | 5 | 0.0015 | 8580 | 24000 | 0.03 | 0.06 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
Direct Connect | 0 | Model Index | 36 | 0 | 05S GP Direct Spans | D | E | F | J | G | H | 0 | IGP05S, Dir, J | IGP05S | 10000 | 1000 | 0 | 1 | 0.075 | 0.075 | 3 | 0.025 | 14300 | 26015 | 0.08 | 0.025 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
Biplanar | 1 | Selections Mismatch | 0 | 0 | 05S GP Biplanar Spans | B | C | D | E | F |  | 0 | IGP05S, Dir, G | IGP05S | 15000 | 5000 | 0 | 1 | 0.2 | 0.2 | 3 | 0.066667 | 19500 | 45000 | 0.08 | 0.025 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
A | 0 | Turndown | 1 | 0 | 05S dP Spans | B | C | D | E |  |  | 0 | IGP05S, Dir, H | IGP05S | 30000 | 10000 | 0 | 1 | 0.2 | 0.2 | 3 | 0.066667 | 33000 | 81000 | 0.08 | 0.025 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
B | 1 | URV Above URL Error | 0 | 0 | 10S AP Direct Spans | D | E |  |  |  |  | 0 | IGP05S, Bi, B | IGP05S | 7.2182 | 0.018046 | -7.2182 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
C | 0 | Pressure Rating Error | 0 | 0 | 10S AP Biplanar Spans | B | C | D | E |  |  | 0 | IGP05S, Bi, C | IGP05S | 36.0912 | 0.089867 | -14.7 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
D | 0 | LRV below LRL Error | 0 | 0 | 10S GP Direct Spans | D | E | F |  |  |  | 0 | IGP05S, Bi, D | IGP05S | 300 | 3.75 | -14.7 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
E | 0 | Small Span Error | 0 | 0 | 10S GP Biplanar Spans | B | C | D | E | F |  | 0 | IGP05S, Bi, E | IGP05S | 3000 | 37.5 | -14.7 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.08 | 0.25 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
F | 0 | Large Span Error | 0 | 1 | 10S dP Spans | B | C | D | E |  |  | 0 | IGP05S, Bi, F | IGP05S | 5000 | 165 | -14.7 | 1 | 0.075 | 0.075 | 5 | 0.015 | 5800 | 14500 | 0.08 | 0.25 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
G | 0 | URV in Reading Range | 1 | 0 | 50S AP Direct Spans | D | E |  |  |  |  | 0 | IDP05S, B | IDP05S | 7.2182 | 0.018046 | -7.2182 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
H | 0 | LRV in Reading Range | 0 | 0 | 50S AP Biplanar Spans | B | C | D | E |  |  | 0 | IDP05S, C | IDP05S | 36.0912 | 0.089867 | -36.0912 | 1 | 0.075 | 0.075 | 5 | 0.015 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
I | 0 | POI in Reading Range | 0 | 0 | 50S GP Direct Spans | D | E | F |  |  |  | 0 | IDP05S, D | IDP05S | 300 | 3.75 | -14.7 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
J | 0 | URV in Ref Range | 1 | 0 | 50S GP Biplanar Spans | B | C | D | E | F |  | 0 | IDP05S, E | IDP05S | 3000 | 37.5 | 0 | 1 | 0.075 | 0.075 | 10 | 0.0075 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.05 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
Transmitter | IDP10S | LRV in Ref Range | 0 | 0 | 50S dP Spans | B | C | D | E |  |  | 0 | IAP10S, Dir, D | IAP10S | 200 | 0.5 | 0 | 25 | 0.05 | 0.05 | 30 | 0.001667 | 300 | 800 | 0.03 | 0.06 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
Analog | 1 | POI in Ref Range | 0 |  | Temp | +/- 50 deg F | No Change |  |  |  |  | 0 | IAP10S, Dir, E | IAP10S | 2000 | 5 | 0 | 25 | 0.05 | 0.05 | 80 | 0.000625 | 3000 | 8000 | 0.03 | 0.06 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Span in Reading Range | 1 |  | Pressure | Up to 1000 PSI | None |  |  |  |  | 0 | IAP10S, Bi, B | IAP10S | 7.2182 | 0.018046 | 0.72182 | 3 | 0.05 | 0.05 | 8 | 0.00625 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Span in Ref Range | 1 |  | RFI | High | None |  |  |  |  | 0 | IAP10S, Bi, C | IAP10S | 36.0912 | 0.089867 | 0 | 3 | 0.05 | 0.05 | 20 | 0.0025 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Temp Factor On | 0 |  | Flow Units | GPM | MGD | SCFH | SCFM | lb/min |  | 0 | IAP10S, Bi, D | IAP10S | 300 | 3.75 | 0 | 25 | 0.05 | 0.05 | 30 | 0.001667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Time in Service | 0 |  | Application | Pressure or Liquid Level | dP Flow |  |  |  |  | 0 | IAP10S, Bi, E | IAP10S | 3000 | 37.5 | 0 | 25 | 0.05 | 0.05 | 10 | 0.005 | 3626 | 14500 | 0.08 | 0.25 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Static Pressure Effect On | 0 |  | Output | Analog (4-20mA) | Digital |  |  |  |  | 0 | IGP10S, Dir, D | IGP10S | 200 | 0.5 | 0 | 25 | 0.05 | 0.05 | 30 | 0.001667 | 300 | 800 | 0.03 | 0.06 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Supply Voltage Shift | 0 | [8 empty] | 0 | IGP10S, Dir, E | IGP10S | 2000 | 5 | 0 | 25 | 0.05 | 0.05 | 80 | 0.000625 | 3000 | 8000 | 0.03 | 0.06 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Vibration | 0 | [8 empty] | 0 | IGP10S, Dir, F | IGP10S | 6000 | 75 | 0 | 3 | 0.05 | 0.05 | 5 | 0.01 | 8580 | 24000 | 0.03 | 0.06 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | RFI On | 0 | [8 empty] | 0 | IGP10S, Bi, B | IGP10S | 7.2182 | 0.018046 | -7.2182 | 3 | 0.05 | 0.05 | 10 | 0.005 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Surge Events | 0 | [8 empty] | 0 | IGP10S, Bi, C | IGP10S | 36.0912 | 0.089867 | -14.7 | 3 | 0.05 | 0.05 | 20 | 0.0025 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | 05S URV Error*Units | 0.1 | [8 empty] | 0 | IGP10S, Bi, D | IGP10S | 300 | 3.75 | -14.7 | 25 | 0.05 | 0.05 | 30 | 0.001667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | 05S LRV Error*Units | 0.1 | [8 empty] | 0 | IGP10S, Bi, E | IGP10S | 3000 | 37.5 | -14.7 | 3 | 0.05 | 0.05 | 3 | 0.01667 | 3626 | 14500 | 0.08 | 0.25 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | 05S POI Error*Units | 0.1 | [8 empty] | 0 | IGP10S, Bi, F | IGP10S | 5000 | 165 | -14.7 | 3 | 0.05 | 0.05 | 5 | 0.01 | 5800 | 14500 | 0.08 | 0.25 | 0 | 0.03 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | 10S/50S URV Error*Units | 0.1 | [8 empty] | 1 | IDP10S, B | IDP10S | 7.2182 | 0.018046 | -7.2182 | 3 | 0.05 | 0.05 | 10 | 0.005 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | 10S/50S LRV Error*Units | 0.033333 | [8 empty] | 0 | IDP10S, C | IDP10S | 36.0912 | 0.089867 | -36.0912 | 3 | 0.05 | 0.05 | 20 | 0.0025 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | 10S/50S POI Error*Units | 0.033333 | [8 empty] | 0 | IDP10S, D | IDP10S | 300 | 3.75 | -14.7 | 10 | 0.05 | 0.05 | 30 | 0.001667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Flow at LRV | 0 | [8 empty] | 0 | IDP10S, E | IDP10S | 3000 | 37.5 | 0 | 10 | 0.05 | 0.05 | 10 | 0.005 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.03 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Flow Error due to Pressure Error at URV | 0.14998 | [8 empty] | 0 | IAP50S, Dir, D | IAP50S | 200 | 0.5 | 0 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 300 | 800 | 0.03 | 0.06 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Flow Error due to Pressure Error at LRV | 8.165 | [8 empty] | 0 | IAP50S, Dir, E | IAP50S | 2000 | 5 | 0 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 3000 | 8000 | 0.03 | 0.06 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Flow Error due to Pressure Error at POI | 0.66228 | [8 empty] | 0 | IAP50S, Bi, B | IAP50S | 7.2182 | 0.018046 | 0.72182 | 4 | 0.025 | 0.025 | 4 | 0.00625 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Totalizing Multiplier | 1440 | [8 empty] | 0 | IAP50S, Bi, C | IAP50S | 36.0912 | 0.089867 | 0 | 5 | 0.025 | 0.025 | 10 | 0.0025 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Totalized URV Error | 4325.395 | [8 empty] | 0 | IAP50S, Bi, D | IAP50S | 300 | 3.75 | 0 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Totalized LRV Error | 12526.0707 | [8 empty] | 0 | IAP50S, Bi, E | IAP50S | 3000 | 37.5 | 0 | 5 | 0.04 | 0.04 | 15 | 0.002667 | 3626 | 14500 | 0.08 | 0.25 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Totalized POI Error | 4424.0156 | [8 empty] | 0 | IGP50S, Dir, D | IGP50S | 200 | 0.5 | 0 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 300 | 800 | 0.03 | 0.06 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
|  | Sig Digits | 1 | [8 empty] | 0 | IGP50S, Dir, E | IGP50S | 2000 | 5 | 0 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 3000 | 8000 | 0.03 | 0.06 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IGP50S, Dir, F | IGP50S | 6000 | 75 | 0 | 5 | 0.04 | 0.04 | 5 | 0.008 | 8580 | 24000 | 0.03 | 0.06 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IGP50S, Bi, B | IGP50S | 7.2182 | 0.018046 | -7.2182 | 5 | 0.025 | 0.025 | 5 | 0.005 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IGP50S, Bi, C | IGP50S | 36.0912 | 0.089867 | -14.7 | 5 | 0.025 | 0.025 | 10 | 0.0025 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IGP50S, Bi, D | IGP50S | 300 | 3.75 | -14.7 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IGP50S, Bi, E | IGP50S | 3000 | 37.5 | -14.7 | 3 | 0.04 | 0.04 | 3 | 0.0133 | 3626 | 14500 | 0.08 | 0.25 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IGP50S, Bi, F | IGP50S | 5000 | 165 | -14.7 | 3 | 0.04 | 0.04 | 4 | 0.01 | 5800 | 14500 | 0.08 | 0.25 | 0 | 0.015 | 0 | 0 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IDP50S, B | IDP50S | 7.2182 | 0.018046 | -7.2182 | 8 | 0.025 | 0.025 | 8 | 0.003125 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IDP50S, C | IDP50S | 36.0912 | 0.089867 | -36.0912 | 5 | 0.025 | 0.025 | 10 | 0.0025 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IDP50S, D | IDP50S | 300 | 3.75 | -14.7 | 5 | 0.025 | 0.025 | 15 | 0.001667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | IDP50S, E | IDP50S | 3000 | 37.5 | 0 | 5 | 0.04 | 0.04 | 15 | 0.002667 | 3626 | 14500 | 0.04 | 0.05 | 0 | 0.015 | 0 | 0.15 | 0 | 0 | 0 | 0.005 | 0 | 0.1 | 0 | 0.1 | 0 | 1 | 0 | 0.01 | 0.02 | 0 | 0
[12 empty] | 0 | Error | Error | Error | Error |  | Error | Error | Error | Error | Error | Error | Error